# Abstract

Validation-first EEG pipelines: how synthetic benchmarks mislead, and how pre-registration catches it

Background: EEG analysis pipelines are usually evaluated on benchmarks their authors design, with outcome measures chosen after the results, and language models summarise such output without checking its numbers.

New method: An open scalp-EEG pipeline (quality control, zero-phase filtering, aperiodic and periodic spectral parameterisation, debiased weighted phase-lag index, a rhythmic-event detector) was built under a validation-first protocol: hypotheses with numeric thresholds were committed before the code or data they concern, methods were tested against synthetic ground truth, and every number in a plain-language summary must cite a computed metric, verified by a rule-based checker.

Results: Four of six initial hypotheses were rejected. A synthetic benchmark on which the detector passed hid four failure modes, including the 1/f confound it targeted; on a held-out subject of the public CHB-MIT database the frozen detector found 4 of 7 seizures at 3.3 false detections per hour. A cloud language model met the citation contract on 30 of 30 recordings and a small local model on 0 of 30, with all faulty text blocked. The first checker missed 709 of 3334 mutated explanations; the corrected one caught all. A blinded audit found 3 of 20 accepted explanations misleading (limit 2): correct numbers attached to the wrong concept.

Comparison with existing methods: Patient-specific seizure detectors reach far higher sensitivity; the contribution is a protocol that makes such claims checkable, not a better detector.

Conclusions: Benchmarks and error models are researcher degrees of freedom; pre-registration and mutation testing make their failures visible.

Keywords: EEG; pre-registration; validation; synthetic benchmark; mutation testing; seizure detection